## Slide 1
Introduction to NGS technologies and library preparation
Chiara Batini, cb334@le.ac.uk
DAY2

## Slide 2
from http://sulab.org/2013/06/sequenced-genomes-per-year/
http://www.ncbi.nlm.nih.gov/
https://gold.jgi.doe.gov/
2004: NHGRI launches ‘$1000 genome’ grants

## Slide 3
platform | # reads | max read length
Roche 454 – 2005 | 200K | 110bp
Illumina (Solexa) GA – 2006 | 30M | 35bp
ABI SOLiD – 2007 | 100M | 35bp
IonTorrent PGM – 2010 | 3M | 100bp
Sequencers: how they were and what they were used for
de-novo assembly
metagenomics
RNA-seq
ChIP-seq
re-sequencing
amplicon-seq

## Slide 5
Sequencers: size and output
Mb
Gb
Tb
bench
unit
Roche 454 GS
Roche 454 GS FLX+
Roche 454 GS Junior
Ion PGM
Ion Proton
Illumina GA
Illumina MiSeq
Illumina NextSeq
Illumina HiSeq
Illumina HiSeq X
Illumina offers:
highest throughput per run
lowest per-base cost
IonTorrent is the fastest

## Slide 6
Sequencers: read length
250-300bp
400bp
1kb
Roche 454 GS FLX+
Illumina MiSeq
Illumina HiSeq
Ion PGM
Ion Proton

## Slide 7
platform | # reads | max read length
Roche | 200K | 110bp
Illumina (Solexa) | 30M | 35bp
ABI SOLiD | 100M | 35bp
IonTorrent | 3M | 100bp
Pacific Biosciences | 5K | up to 4kb
# reads | max read length
will be discontinued in 2016 | 
5BN | 300bp
will be/is discontinued…? | 
80M | 400bp
50K | up to 20kb
Sequencers: how they are
250-300bp
400bp
1kb
>10kb

## Slide 8
PacBio
Nanopore MinION
250-300bp
400bp
1kb
Roche 454 GS FLX+
Illumina
MiSeq
Illumina HiSeq
Ion PGM
Ion Proton
>10kb
Sequencers: the real-time generation
Pacific Biosciences (PacBio) - 2010
Nanopore MinION - 2014

## Slide 9
Sequencers: size and output
Mb
Gb
Tb
bench
unit
room
Roche 454 GS
Roche 454 GS FLX+
Roche 454 GS Junior
Ion PGM
Ion Proton
Illumina GA
Illumina MiSeq
Illumina NextSeq
Illumina HiSeq
Illumina HiSeq X
PacBio
pocket
Nanopore MinION
Nanopore PromethION

## Slide 10
Illumina sequencing technology (1)
Sequencing by synthesis, 4 modified dNTPs together, imaging
support: slide
from https://www.illumina.com/documents/products/techspotlights/techspotlight_sequencing.pdf

## Slide 11
Illumina sequencing technology (2)
from https://www.illumina.com/documents/products/techspotlights/techspotlight_sequencing.pdf

## Slide 12
Illumina sequencing technology (3)
from https://www.illumina.com/documents/products/techspotlights/techspotlight_sequencing.pdf

## Slide 13
IonTorrent sequencing technology (1)
from https://www.thermofisher.com/uk/en/home/life-science/sequencing/next-generation-sequencing/ion-torrent-next-generation-sequencing-technology.html
Sequencing by synthesis, 1 non-modified dNTP at the time, pH detection
support: bead/chip

## Slide 14
IonTorrent sequencing technology (2)
from https://www.thermofisher.com/uk/en/home/life-science/sequencing/next-generation-sequencing/ion-torrent-next-generation-sequencing-technology.html

## Slide 15
Sequencers: error rates and patterns
Instrument | Primary Errors | Single-pass Error Rate (%) | Final Error Rate (%)
ABI 3730xl (capillary) | substitution | 0.1-1 | 0.1-1
Roche 454 – All models | indel | 1 | 1
Illumina – All models | substitution | ~0.1 | ~0.1
Ion Torrent – all chips | Indel | ~1 | ~1
Oxford Nanopore | deletions | ≥4* | 4*
PacBio RS | Indel | ~13 | ≤1
from http://www.molecularecologist.com/next-gen-fieldguide-2014/

## Slide 16
NGS applications
Genomic DNA:
WGS (whole genome sequencing): population-scale; single-cell
WES (whole exome sequencing)
custom enrichment
amplicon-seq
RAD-seq (restriction-site-associated DNA): very useful for organisms lacking reference genomes

## Slide 17: De novo assembly

## Slide 18: Re-sequencing
Reference genome
Map reads to reference
Call Variants
SNP
SNP
SNP
SNP

## Slide 19: What does DNA-seq tell you
Calling of known polymorphisms
Identification of novel polymorphism (SNPs, indels etc)
Genomic rearrangements, large deletions and insertions
What the genome of your sample looks like (genome assembly, chromosome assembly)

## Slide 20
RNA-seq:
strand-specific RNA-seq protocols
single-cell transcriptomics
Fluorescent in situ RNA seq (FISSEQ): it allows the characterization of the single-cell transcriptomics and the localization of the transcripts withi the cell
CaptureSeq
Native elongating transcript seq (NET-seq)
NGS applications

## Slide 21: RNA-seq

## Slide 22: ChIP-seq
Location-based techniques
ChIP-seq: protein-DNA interactions; ChIP-exo (down to nucleotide level)
protein-RNA, RNA-DNA, DNA-DNA interactions

## Slide 23
from http://www.mdpi.com/2079-7737/2/1/378/htm
Diagnostic tools in clinical laboratories; spreading in forensics
Illumina HiSeq X Ten : a human genome for less than $1000?
Personalized medicine is just around the corner…

## Slide 24
So, which is the best technology?
https://docs.google.com/spreadsheets/d/1eoWVtwKvnbV8MaRb9hzCSpGtEjazO0XTt3a8TbPG4dY/edit?pli=1#gid=0
Good support and communication are essential

## Slide 25
So, how do we get our DNA ready to be sequenced then?

## Slide 26
DNA source and quantity
library prep
clonal amplification
sequencing

## Slide 27
DNA source and quantity
whole genome 2ug for paired end, 20ug for mate pair
Exome 6ug or more
custom enrichment min 5ug, more depending on the target size
amplicon seq a few ng per amplicon may be enough, since PCR will increase the concentration
non-degraded, RNA-free, high quality DNA
quality can be checked with nanodrop, specific ratios are suggested:
260/280: 1.7-2.0
260/230: >2
for quantification other methods are suggested [Qubit, picogreen]

## Slide 28
DNA source and quantity
library prep
shearing
adapter ligation and barcoding
size selection
amplification
quantification
whole genome
exome
custom enrichment
amplicon seq
RNA
sequencing
clonal amplification

## Slide 29
library prep - shearing
Two main methods are used:
Sonication: hydrodynamic shearing using acoustic energy; bubbles are formed in solution, when they explode they break the DNA
Enzymatic reaction: enzymes randomly cut the dsDNA in a time-dependent manner
modified from Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402

## Slide 30
library prep – adapter ligation and barcoding
adapters: 30-50bp fragments which contain primer sites for amplification and are needed to link the fragment with the support (slide, bead)
barcodes/indexes: 6-10bp fragments which carry a unique sequence; they are used to distinguish samples run in the same lane/chip
Adapters and barcodes are combined into one fragment
modified from Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402

## Slide 31
library prep – size selection
Needed to select fragments of the right length from a mixture of fragments generated by the fragmentation step: it is usually 200-300bp for paired end sequencing with Illumina, but it may vary depending on the read length for 454 and IonTorrent

## Slide 32
Amplification: needed to increase the concentration of the fragments which positively incorporated adapters
library prep – amplification
modified from Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402

## Slide 33
Quantification: needed to tune the quantity of template for the run – usually performed with the Agilent Bioanalyzer
library prep –quantification

## Slide 34
DNA source and quantity
library prep
shearing
adapter ligation and barcoding
size selection
amplification
quantification
clonal amplification
ePCR (IonTorrent/454)
clustering (Illumina)
whole genome
exome
custom enrichment
amplicon seq
RNA
sequencing

## Slide 35
clonal amplification – ePCR (Roche 454/Ion Torrent)
emulsion PCR: ideally each drop contains a single fragment, PCR reagents and a bead with primers [beads may happen to be polyclonal, will be discarded after sequencing]
each bead will have one million copies of each fragment on its surface
an enrichment step is needed afterwards in order to increase the concentration of beads with attached fragments
modified from Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402

## Slide 36
clonal amplification – clustering Illumina
isothermal amplification: temperature is kept constant but reagent are cycled to perform amplification, not very efficient (on purpose)
modified from Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402

## Slide 37
clusters usually contain around 1000 identical copies of a single template
optimal cluster density will provide optimal sequencing output
clonal amplification – clustering Illumina
modified from Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402

## Slide 38
A few videos…
Illumina
https://www.youtube.com/watch?v=womKfikWlxM
IonTorrent
https://www.youtube.com/watch?v=ZL7DXFPz8rU
PacBio
https://www.youtube.com/watch?v=NHCJ8PtYCFc
Oxford Nanopore
https://www.youtube.com/watch?v=3UHw22hBpAk

## Slide 39
useful references and links
Mardis 2008 Annu Rev Genomics Hum Genet 9:387-402
Metzker 2010 Nat Rev Genet 11:31-46
http://seqanswers.com/forums/index.php
Thanks to:
Pierpaolo Maisano Delser
Rita Neumann
James Eales
http://www.frontlinegenomics.com/1649/next-generation-sequencing-how-and-why-we-got-here/
http://thewestheimerinstitute.org/pubs/The%20challenges%20of%20sequencing%20by%20synthesis.pdf
http://www.molecularecologist.com/next-gen-fieldguide-2014/

## Slide 40
Data QC
Mapping
Local Realignment
Base quality recalibration
Duplicate removal
Variant Calling

## Slide 41
Data QC
Mapping
Local Realignment
Base quality recalibration
Duplicate removal
Variant Calling